5. PROJECT NUMBER (if applicable)

CODE

7. ADMINISTERED BY

2. AMENDMENT/MODIFICATION NUMBER

CODE

6. ISSUED BY

8. NAME AND ADDRESS OF CONTRACTOR

4. REQUISITION/PURCHASE REQ. NUMBER

3. EFFECTIVE DATE

9A. AMENDMENT OF SOLICITATION NUMBER

9B. DATED

PAGE        OF     PAGES

10A. MODIFICATION OF CONTRACT/ORDER NUMBER

10B. DATED

BPA NO.

1. CONTRACT ID CODE

FACILITY CODE

CODE

Offers must acknowledge receipt of this amendment prior to the hour and date specified in the solicitation or as amended, by one of the following methods:

The above numbered solicitation is amended as set forth in Item 14.  The hour and date specified for receipt of Offers

E. IMPORTANT:

is extended,

(a) By completing Items 8 and 15, and returning __________ copies of the amendment; (b) By acknowledging receipt of this amendment on each copy of the

offer submitted; or (c) By separate letter or electronic communication which includes a reference to the solicitation and amendment numbers.  FAILURE OF YOUR

ACKNOWLEDGMENT TO BE RECEIVED AT THE PLACE DESIGNATED FOR THE RECEIPT OF OFFERS PRIOR TO THE HOUR AND DATE SPECIFIED MAY

is not extended.

12. ACCOUNTING AND APPROPRIATION DATA

(REV. 11/2016)

is required to sign this document and return ___________ copies to the issuing office.

is not,

A. THIS CHANGE ORDER IS ISSUED PURSUANT TO:    (Specify authority)   THE CHANGES SET FORTH IN ITEM 14 ARE MADE IN THE CONTRACT ORDER NO. IN ITEM 10A.

15C. DATE SIGNED

B. THE ABOVE NUMBERED CONTRACT/ORDER IS MODIFIED TO REFLECT THE ADMINISTRATIVE CHANGES

SET FORTH IN ITEM 14, PURSUANT TO THE AUTHORITY OF FAR 43.103(b).

RESULT IN REJECTION OF YOUR OFFER.  If by virtue of this amendment you desire to change an offer already submitted, such change may be made by letter

or electronic communication, provided each letter or electronic communication makes reference to the solicitation and this amendment, and is received prior to

the opening hour and date specified.

C. THIS SUPPLEMENTAL AGREEMENT IS ENTERED INTO PURSUANT TO AUTHORITY OF:

D. OTHER

Contractor

16C. DATE SIGNED

14. DESCRIPTION OF AMENDMENT/MODIFICATION

16B. UNITED STATES OF AMERICA

Except as provided herein, all terms and conditions of the document referenced in Item 9A or 10A, as heretofore changed, remains unchanged and in full force and effect.

15A. NAME AND TITLE OF SIGNER

16A. NAME AND TITLE OF CONTRACTING OFFICER

15B. CONTRACTOR/OFFEROR

STANDARD FORM 30

PREVIOUS EDITION NOT USABLE

Prescribed by GSA - FAR (48 CFR) 53.243

(Type or print)

(Type or print)

(Organized by UCF section headings, including solicitation/contract subject matter where feasible.)

(Number, street, county, State and ZIP Code)

(If other than Item 6)

(Specify type of modification and authority)

(such as changes in paying office, appropriation date, etc.)

(If required)

(SEE ITEM 11)

(SEE ITEM 13)

(X)

CHECK

ONE

13. THIS ITEM APPLIES ONLY TO MODIFICATIONS OF CONTRACTS/ORDERS,

IT MODIFIES THE CONTRACT/ORDER NO. AS DESCRIBED IN ITEM 14.

11. THIS ITEM ONLY APPLIES TO AMENDMENTS OF SOLICITATIONS

AMENDMENT OF SOLICITATION/MODIFICATION OF CONTRACT

(Signature of person authorized to sign)

(Signature of Contracting Officer)

1

11

0001

09-08-2025

None

36C263

Department of Veterans Affairs

NETWORK 23 CONTRACTING OFFICE

Contract Officer

4801 Veterans Drive

St. Cloud MN 56303

Traci Johannes, Contracting Officer

Same as Block 6: ISSUED BY

To all Offerors/Bidders

36C26325Q1095

X

X

X

X

1

The purpose of this solicitation amendment is to address the following:

1. Attach Statement of Work again due to technology error.

Traci Johannes

Contracting Officer

## Statement of Work

## VISN 23 Grossing Station Replacement Procurement

- Background:  Several sites in VISN 23 are in need of replacing their current grossing stations in the Histology Laboratories. These grossing stations have reached the end of their life expectancy and we have begun to see an increase in the stations experiencing issues that require a service call. The increased downtime of these machines greatly impacts the speed and efficiency of the labs throughout the VISN in performing specimen examination. Additionally, the labs are seeking a newer model of grossing station which has additional features that would increase the speed, efficiency, and safety of the staff performing specimen examinations.

- Purpose of Project:  Replace 8 grossing stations throughout VISN 23. Specifically, 4 in Minneapolis and 1 each in Omaha (NWI), Fargo, Central Iowa, and Iowa City. Each of these sites have grossing stations that are past their life expectancy and are experiencing an increase in service calls.

- Statement of Work

  - Description of Work:

The vendor will provide the following quantities of grossing stations with the following BRAND NAME OR EQUAL Maestro Encore features as required by site:

Minneapolis

QTY 4 Grossing Stations

- No larger than 60 inches by 30 inches by 80 inches to fit into existing space

- Right hand sink

- Air Flow Detection with backdraft alert monitoring system

- Adjustable lighting

- Ability to direct air from the fron to the back of the cutting surface

- Integrated Air Filtration System to prevent dust & debris build up

- Backdraft ventilation

- Dual 8″ duct stubs for in-house ventilation

- Hot/Cold Hand Sprayer

- Ergorail™ Track for mounting of monitor & keyboard arms

- Pegboard for mounting of accessories

  - Include dye bottle holder

  - Include large shelf at least 14” L x 8” D x 2” H

  - Include small shelf at least 7” L x 8” D x 2” H

  - Include bin for papertowels/c-fold gloves

  - Include magnetic toolbar

- Two (2) GFCI and two (2) USB receptacles, one (1) each with continuous power

- Sliding work surface with interchangeable workflow components

- One (1) sliding work surface perforated grid plate

- One (1) sliding work surface dissecting board, white

- Two (2) wire ports for cord management

- 1/2 Hp. disposal

- Magnetic instrument bar

- One (1) large and one (1) small organizer bin

- Seismic anchoring points

- Touchscreen monitor with automated image enhancement

- Monitor/keyboard mounting column

- Dual gooseneck LED task light

- At least 3 (2 small – at least 6” W x W 7” D x 4” H and 1 large – at least 26.5” W x 7” D X 4” H)) drawers per grossing station

- Rinsing system for work surface with temperature and flow control

- Formalin collection system

- Faucet with foot pedal control

Omaha (NWI)

QTY 1 Grossing Station

- No larger than 60 inches by 30 inches by 80 inches to fit into existing space

- Right hand sink

- Air Flow Detection with backdraft alert monitoring system

- Adjustable lighting

- Ability to direct air from the fron to the back of the cutting surface

- Integrated Air Filtration System to prevent dust & debris build up

- Backdraft ventilation

- Dual 8″ duct stubs for in-house ventilation

- Hot/Cold Hand Sprayer

- Ergorail™ Track for mounting of monitor & keyboard arms

- Pegboard for mounting of accessories

  - Include large shelf at least 14” L x 8” D x 2” H

- Two (2) GFCI and two (2) USB receptacles, one (1) each with continuous power

- Sliding work surface with interchangeable workflow components

- One (1) sliding work surface perforated grid plate

- One (1) sliding work surface dissecting board, white

- Two (2) wire ports for cord management

- 1/2 Hp. disposal

- Magnetic instrument bar

- One (1) large and one (1) small organizer bin

- Seismic anchoring points

- Touchscreen monitor with automated image enhancement

- Formalin collection system

- Hands free infared faucet

- Cabinet doors

Fargo

QTY 1 Grossing Station

- No larger than 60 inches by 30 inches by 80 inches to fit into existing space

- Right hand sink

- Air Flow Detection with backdraft alert monitoring system

- Adjustable lighting

- Ability to direct air from the front to the back of the cutting surface

- Integrated Air Filtration System to prevent dust & debris build up

- Backdraft ventilation

- Dual 8″ duct stubs for in-house ventilation

- Hot/Cold Hand Sprayer

- Ergorail™ Track for mounting of monitor & keyboard arms

- Pegboard for mounting of accessories

  - Include dye bottle holder

  - Include large shelf at least 14” L x 8” D x 2” H

  - Include small shelf at least 7” L x 8” D x 2” H

  - Include bin for papertowels/c-fold gloves

  - Include requisition holder – should fit 8.5” x 11”requisitions

- Two (2) GFCI and two (2) USB receptacles, one (1) each with continuous power

- Sliding work surface with interchangeable workflow components

- One (1) sliding work surface perforated grid plate

- One (1) sliding work surface dissecting board, white

- Two (2) wire ports for cord management

- 1/2 Hp. disposal

- Magnetic instrument bar

- One (1) large and one (1) small organizer bin

- Seismic anchoring points

- Touchscreen monitor with automated image enhancement

- Monitor/keyboard mounting column

- Single gooseneck LED task light

- At least 2 small – at least 6” W x W 7” D x 4” H drawer

- Rinsing system for work surface with temperature and flow control

- Formalin collection system

- Faucet with foot pedal control

Central Iowa

QTY 1 Grossing Station

- no larger than 60 inches by 30 inches by 80 inches to fit into existing space

- Right hand sink

- Air Flow Detection with backdraft alert monitoring system

- Adjustable lighting

- Ability to direct air from the fron to the back of the cutting surface

- Integrated Air Filtration System to prevent dust & debris build up

- Backdraft ventilation

- Dual 8″ duct stubs for in-house ventilation

- Hot/Cold Hand Sprayer

- Ergorail™ Track for mounting of monitor & keyboard arms

- Pegboard for mounting of accessories

  - Include bin for papertowels/c-fold gloves

  - Include requisition holder – should fit 8.5” x 11”requisitions

- Two (2) GFCI and two (2) USB receptacles, one (1) each with continuous power

- Sliding work surface with interchangeable workflow components

- One (1) sliding work surface perforated grid plate

- One (1) sliding work surface dissecting board, white

- Two (2) wire ports for cord management

- 1/2 Hp. disposal

- Magnetic instrument bar

- One (1) large and one (1) small organizer bin

- Seismic anchoring points

- Touchscreen monitor with automated image enhancement

- Keyboard arm mounting column

- Single gooseneck LED task light

- At least 2 small – at least 6” W x W 7” D x 4” H drawer

- Rinsing system for work surface with temperature and flow control

- Formalin collection system

- Under shelf lighting for work space

Iowa City

QTY 1 Grossing Station

- no larger than 72” inches by 30 inches by 80 inches to fit into existing space

- Right hand sink

- Air Flow Detection with backdraft alert monitoring system

- Adjustable lighting

- Ability to direct air from the fron to the back of the cutting surface

- Integrated Air Filtration System to prevent dust & debris build up

- Backdraft ventilation

- Dual 8″ duct stubs for in-house ventilation

- Hot/Cold Hand Sprayer

- Ergorail™ Track for mounting of monitor & keyboard arms

- Pegboard for mounting of accessories

  - Include 2 dye bottle holders

  - Include face shield holster

  - Include bin for papertowels/c-fold gloves

  - Include requisition holder – should fit 8.5” x 11”requisitions

  - Include 2 large shelves at least 14” L x 8” D x 2” H

  - Include 3 small shelves at least 7” L x 8” D x 2” H

  - Include magnetic toolbar

- Two (2) GFCI and two (2) USB receptacles, one (1) each with continuous power

- Sliding work surface with interchangeable workflow components

- One (1) sliding work surface perforated grid plate

- One (1) sliding work surface dissecting board, white

- Two (2) wire ports for cord management

- 1/2 Hp. disposal

- Magnetic instrument bar

- One (1) large and one (1) small organizer bin

- Seismic anchoring points

- Touchscreen monitor with automated image enhancement

- Keyboard/monitor mounting column

- CPU mount

- dual gooseneck LED task light

- At least 3 (1 small – at least 6” W x W 7” D x 4” H drawer, 1 small instrument – at least 6” W x W 7” D x 4” H with drain petcock, and 1 large - at least 26.5” W x 7” D X 4” H) drawers

- Rinsing system for work surface with temperature and flow control

- Under shelf lighting for work space

- Faucet with foot pedal

- Cabinet doors

Vendor is to provide a copy of equipment operator documentation either in paper or electronic format.

The equipment models being offered shall be in current production as of the date this offer is submitted.  For purposes of this solicitation, “current product” shall mean that the grossing station is being offered as new equipment.  Discontinued models that are only being made available as remanufactured equipment are not acceptable.  Refurbished, used or current in house equipment is not acceptable.

- Equipment Warranty

All of the equipment should come with a 1 year warranty and include any and all preventative maintenance services recommended by the original equipment manufacturer during that warranty period.

- Deinstallation of old equipment and Installation of New Equipment

The vendor shall remove 3 existing grossing stations at  Minneapolis and 1 existing grossing station each at Omaha (NWI), Fargo, Central Iowa, and Iowa City. The lab will be responsible for decontaminating the grossing station and the vendor shall dismantle and take the equipment to the loading dock. The VA will be responsible for disposable of the equipment. The vendor will be responsible for fully installing the equipment and then testing for correct function. The vendor will ensure the grossing station air flow/ventilation is functioning properly after installation by having it certified.

- Training

The vendor shall provide at least 1 day of onsite training for up to 8 people. This should be offered over at least 2 different shifts if needed by the lab.

- Integrations

Must integrate with each facilities HVAC system. No other integrations – this equipment will not be connected to the network.

- Schedule of Deliverables:

All equipment and supplies shall be delivered to the following VA medical centers.  Equipment shall be delivered within 60 days of receipt of an order.

Deliveries shall be made to:

VA Central Iowa Health Care System

3600 30th St, Des Moines, IA 50310

VA Iowa City Health Care System

601 Highway 6 West, Iowa City, IA 52246-2209

VA Minneapolis Health Care System

1 Veterans Dr, Minneapolis, MN 55417

VA Nebraska-Western Iowa Health Care System

4101 Woolworth Ave, Omaha, NE 68105

VA Fargo Health Care System

2101 Elm St NE, Fargo, ND 58102

Delivery times: Between 8:00AM and 3:30PM, Monday-Friday except Federal Holidays.   All deliveries shall be accompanied by a delivery ticket or sales slip which shall contain:

Vendor’s name

Purchase order number

Date of order

Date of delivery

Itemized list of products furnished including product description and quantity shipped.

- Certificates of Performance

Company shall provide “certificates of performance” shipped stating compliance with your company’s standards and CLSI (formerly NCCLS) standards where they exist.

- Recalled and Defective Products

The contractor will immediately notify the Contracting Officer of any recalls of product or other important safety issues.  As appropriate, the Contractor will replace and/or reimburse recalled/defective products at no cost to the Government.  The Contractor may be liable for costs of processing recalls, i.e. administrative and clinical services to replace recalled/defective products.