## Sheet1
Covered Equipment
| Vendor functional Code | Serial Numbers | Equipment Description | EE #
| FLOC's: 0022431842, 0022431843 | 1369829, 1371816 | CWP 102H System | NA, NA
| FLOCS's: 0022431834, 0022431835, | 1381384, 1381385, 1381386, | MHX Portables (7) | 114977, 114975,114976
| 0022431836, 0022431837, 0022431838, | 1381387, 1383780, 1383783, |  | 114974, 116275, 116274
| 0022431839, 0022431840 | 1383876 |  | 116276
Schedule of Services
2 | System Performance Review (W3TSP7312) |  |  | SI0002 | Monthly | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2
4 | Rebuild Tank Head (2850,NHWBP) (W3TSP7034:includes: W2T132346-1, W2T913251-1, Labor 1 HR) |  |  | SRV1001 | Annually | 4
4 | Carbon Filter Rebed, 7 cu.ft. (W3T572487:includes:W3T572487-7, W2T911945-4, W3T577797-1, W2T913121-1, W2T913131-1, LABOR-3 HR) |  |  | SRV1010 | Annually | 4
2 | Clean Membranes (W3TSP4638:includes: W2T934418 -.25, W2T912365-.25, LABOR (4 HR)) |  |  | 4000-0215 | Semi-Annually | 2 | [5 empty] | 2
2 | CWP Annual Preventative Maintenance Kit (W3T577722-every 12 months, W3T577723-every 24 months, W3T577669-every 48 months) |  |  | 4000-0092 | Annually | 2
2 | Inspect & Replace RO Pre Filter (W3TSP4603) |  |  | SI5000 | Monthly | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2
2 | FILTER,DEPTH,SW,PP,1M,4-1/2"X20",DOE (W2T913146) |  |  | ME42007 | Monthly | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2
2 | Inspect & Replace Post Filter (W3TSP4602) |  |  | SI5120 | Annually | 2
2 | Filter, RenaPure, 20", 222, FLAT (W3T590100) |  |  | W3T590100 | Annually | 2
4 | Rental DI Tanks, 14x47, 3.6 cu.ft. (W5TDIMB10360FSQR) |  |  | 920-01-047 | Semi-Annually | 4 | [5 empty] | 4
1 | AAMI (W3TSP7119)Test City Water (Customer to Pull) |  |  | AAMI TEST | Annually | 1
2 | AAMI (W3TSP7119) RO Product (Customer to Pull) |  |  | AAMI TEST | Semi-Annually | 2 | [5 empty] | 2
90 | LAL Water Test (W3TSP7171) (Customer to Pull) |  |  | LAL TEST | Monthly | 90 | 90 | 90 | 90 | 90 | 90 | 90 | 90 | 90 | 90 | 90 | 90
90 | Bacteria Water Test (W3TSP7170) (Customer to Pull) |  |  | BACTERIA TEST | Monthly | 90 | 90 | 90 | 90 | 90 | 90 | 90 | 90 | 90 | 90 | 90 | 90
6 | LAL Water Test (W3TSP7171) (Evoqua to Pull) |  |  | LAL TEST | Monthly | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6
6 | Bacteria Water Test (W3TSP7170) (Evoqua to Pull) |  |  | BACTERIA TEST | Monthly | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6 | 6
1 | LEAVE GALLON MINNCARE HD AS NEEDED (W2T912234) |  |  | 3029764 | Monthly | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
1 | TEST,STRIP,MINNCARE HD,RESID,100 (W2T912238 ) |  |  | 3029795 | Quarterly | 1 |  |  | 1 |  |  | 1 |  |  | 1
2 | CWP Extended Coverage (W3TSP3048) |  |  | 4000-1092 | Annually | 2
1 | Annual Cleaning of Temperature Blending Valve (W3TSP4539 ) |  |  | 4000-1101 | Annually | 1
1 | Annual Calibration of Booster Pump (W3TSP7356) |  |  | 4000-1102 | Annually | 1
7 | Portable Performance Review (W3TSP7037) |  |  | 4000-0089 | Monthly | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7
7 | Portable Chemical Disinfection (Multiple) (W3TSP4707:includes: W2T912235-.1, W2T912238-.1, W2T912239-.25, LABOR-1 HR) |  |  | 4000-0083 | Quarterly | 7 |  |  | 7 |  |  | 7 |  |  | 7
7 | Portable Cleaning (High and Low) (W3TSP4638: includes: W2T934418-.1, W2T912365-.1, Labor-2hr) |  |  | 4000-0060 | Quarterly | 7 |  |  | 7 |  |  | 7 |  |  | 7
7 | Inspect & Replace RO Pre Filter (W3TSP4603) |  |  | SI5000 | Monthly | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7
7 | FILTER,5M,9-7/8IN (W2T913469) |  |  | ME40696 | Monthly | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7
7 | Inspect & Replace 1 Carbon Block PTG-520® Filter (W3TSP4763) |  |  | 4000-0073 | Quarterly | 7 |  |  | 7 |  |  | 7 |  |  | 7
7 | Carbon Block PTG (W2T912087) |  |  | 3027145 | Quarterly | 7 |  |  | 7 |  |  | 7 |  |  | 7
7 | AAMI (W3TSP7119) |  |  | AAMI TEST | Semi-Annually | 7 | [5 empty] | 7
7 | LAL Water Test (48 Hour) (W3TSP7171) (Evoqua to Pull) |  |  | LAL TEST | Monthly | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7
7 | Bacteria Water Test (W3TSP7170) (Evoqua to Pull) |  |  | BACTERIA TEST | Monthly | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7
1 | Unscheduled/Emergency (W3TSP7001) |  |  | (W3TSP7001)
| RO system operation verification |  |  |  | quarterly
| Carbon tank rebed x 2 |  |  |  | annually
| RO unit - storage tank vent filters |  |  |  | semi annually
| RO unit- control panel resistivity monitor |  |  |  | annually
| UV sterilizer - bulb replacement |  |  |  | annually
| DI mixed bed polishing tanks x 2 |  |  |  | quarterly
| final filter endotoxin replacement |  |  |  | semi annually
| disinfection of central systems (after hours, Sunday) |  |  |  | monthly
unscheduled/emergency work on RO system