Notice ID Number:  NIST-SS26-CHIPS-23

Title: Mid-Infrared Camera

Response Date:  December 15,2025

The purpose of this sources sought notice is to conduct market research and identify potential sources of commercial products/services that satisfy the Government’s anticipated needs.

BACKGROUND

The Digital Twins for Atomic Layer Deposition Processes Project (5.05) was undertaken to develop and validate digital twin (DT) models for atomic layer deposition (ALD) processes to help facilitate efficient design and operation of these processes. One aspect of this project involves development and application of in situ optical diagnostics capable of quantifying precursor chemical concentrations in the deposition chamber, in both one and two dimensions. One optical technique that will be employed is mid-infrared (IR) gas-phase absorption. For one implementation of this technique, Group 646.05 must procure a mid-IR camera to record images of molecular absorptions by precursor chemicals in a deposition chamber. Since the precursor chemicals will be present at low concentrations, it is critical to have a camera with high quantum efficiency and low noise. Furthermore, since different precursor compounds will be investigated, it is critical to have a good response over a wide spectral range. In addition, since the flow system is dynamic, a camera capable of high frame rates is needed. Also, the lens must have characteristics that permit operation with a large depth of field so that as much of the chamber volume is in focus as possible. Hence, the lens characteristics must be well matched to the camera design. Additionally, 646.05 will be making measurements in two mid-IR wavelength ranges, the mid-wave IR and long-wave IR spectral regions, and, therefore, 646.05 needs to purchase separate lenses optimized for operation at each wavelength range.

The purpose is to purchase a mid-infrared camera and lenses suitable for high-sensitivity, large depth of field, and high-speed absorption imaging measurements across the mid-IR spectral region.

NIST is seeking information from sources that may be capable of providing a solution that will achieve the objectives described above, in addition to the following essential requirements:

TECHNICAL REQUIREMENTS:

ENTER REQUIREMENTS

Line Item 0001: Mid-IR camera

Description:  Mid-IR camera capable of high-sensitivity and high-speed measurements across the mid-infrared spectral region.

Quantity:  One (1)

- Technical Specifications

  - Wavelength range of operation:  a wavelength of 3 microns or shorter to a wavelength of 10.5 microns or longer

  - Sensor type:  quantum (e.g., not microbolometer or thermopile)

  - Sensor architecture:  strained layer superlattice (SLS)

  - Pixel Format:  greater than or equal to 640 pixels by greater than or equal to 512 pixels

  - Pixel Dimensions:  less than or equal to 25 microns by less than or equal to 25 microns (square)

  - Pixel well capacity:  Greater than or equal to 5 million electrons

  - Frame readout rage (full array):  greater than or equal to 200 frames per second

  - Noise equivalent change in temperature:  45 mK or less

  - Digitization:  greater than or equal to 14 bits

  - f-number:  between f/2.0 and f/4.0

  - Filter wheel with at least four filter positions (one blank and three filters) located in the cold section of the camera

  - Filters must be changeable by the user (MML/646.05)

  - Filter diameter:  25.4 mm

  - Mounting:  Infrared Bayonet mount and 4x #6-32 screws on a 2.63 inch bolt hole)

Line Item 0002:  Mid-wave IR lens

Description:  The characteristics of this lens must be matched to the camera and this lens must be capable of operation in the mid-wave IR spectral region.

Quantity: One (1)

- Technical Specifications

  - Wavelength range of operation:  a wavelength of 3 microns or shorter to a wavelength of 5 microns or longer

  - In-band Transmission:  >90%

  - Focal length:  50 mm

  - f-number:  between f/2.0 and f/4.0

  - Minimum object distance:  500 mm or less

  - Image diagonal:  21 mm

  - Circular field of view:  23 degrees or greater

  - Stop location:  25.4 mm from sensor (must match cold stop position of camera)

  - Stop diameter:  11 mm (must match the diameter of the camera cold stop)

  - Back Working Distance:  33.1 mm (front of camera housing to sensor)

  - Mounting:  must be compatible with the camera (possibly either Infrared Bayonet mount or 4x #6-32 screws on a 2.63 inch bolt hole)

  - Focus:  manual

Line Item 0003:  Long-wave IR lens

Description:  The characteristics of this lens must be matched to the camera and this lens must be capable of operation in the long-wave IR spectral region.

Quantity: One (1)

- Technical Specifications

  - Wavelength range of operation:  a wavelength of 8 microns or shorter to a wavelength of 12 microns or longer

  - In-band Transmission:  >90%

  - Focal length:  50 mm

  - f-number:  between f/2.0 and f/4.0

  - Minimum object distance:  500 mm or less

  - Image diagonal:  21 mm

  - Circular field of view:  23 degrees or greater

  - Stop location:  25.4 mm from sensor (must match cold stop position of camera)

  - Stop diameter:  12.4 mm (must match the diameter of the camera cold stop)

  - Back Working Distance:  33.1 mm (front of camera housing to sensor)

  - Mounting:  must be compatible with the camera (possibly either Infrared Bayonet mount or 4x #6-32 screws on a 2.63 inch bolt hole)

  - Focus:  manual

HOW TO RESPOND TO THIS NOTICE

In responding to this notice, please DO NOT PROVIDE PROPRIETARY INFORMATION. Please include only the following information, readable in either Microsoft Word 365, Microsoft Excel 365, or .pdf format, in the response:  Submit the response by email to the Primary Point of Contact and, if specified, to the Secondary Point of Contact listed in this notice as soon as possible, and preferably before the closing date and time of this notice.  Please note that to be considered for award under any official solicitation, the entity must be registered and “active” in SAM at the time of solicitation response.

- Provide the complete name of your company, address, name of contact for follow-up questions, their email, their phone number and, if your company has an active registration in https://sam.gov, your company’s Unique Entity ID (UEI).

- Identify the equipment that your company sells that can meet the objectives addressed in the BACKGROUND section of this notice.  For each product recommended to meet the Government’s requirement, provide the following:

  - Manufacturer name

  - Model number

  - Technical specifications

  - If your company is not the manufacturer, provide information on your company’s status as an authorized reseller of the product(s)

- Describe performance capabilities of the product(s) your company recommends to meet the Government’s requirements.  Additionally, if there are other features or functions that you believe would assist NIST in meeting its objectives described above, please discuss in your response.

  - Discuss whether the equipment that your company sells and which you describe in your response to this notice may be customized to specifications and indicate any limits to customization.

- Identify any aspects of the NIST market research notice, including instructions, and draft minimum specifications in the BACKGROUND section you cannot meet and state why.  Please offer suggestions for how the market research notice and draft minimum specifications could be made more competitive.

- State whether the proposed equipment is manufactured in the United States and, if not, state the name of the country where the equipment is manufactured.

- Identify any plans/possibilities for changes in manufacturing location of the aforementioned equipment and provide relevant details, including timeline.

- For the NAICS code listed in this notice:

  - Indicate whether your company is (a) a small business or (b) other than small business.  See the Table of Small Business Size Standards and the associated .pdf download file for small business size standards and additional information.

  - If you believe the NAICS code listed in this notice is not the best NAICS code for the type of product addressed in this notice, identify an alternative NAICS code that you believe would be more appropriate for the planned procurement.

- Describe services that are available with the purchase of the aforementioned equipment from your company such as installation, training, and equipment maintenance.

- Describe standard terms and conditions of sale offered by your company for the recommended equipment such as: delivery time after your company accepts the order; FOB shipping terms; manufacturer warranty (including details regarding nature and duration); if available, description(s) of available extended warranty; equipment setup and test; operator and service instruction manual(s); cleanup after installation; and if applicable, other offered services.  Provide a copy of manufacturer standard terms and conditions that typically relate to the sale of the specified equipment, if available.

- State whether your company offers facility renovation services related to installation of the recommended equipment at its delivery destination, if required per the NIST-identified minimum specifications, and provide description of said services.  Indicate if your company performs the facility renovation services or typically subcontracts the work to another company.  Indicate if your company would be interested in inspecting the intended installation site during the market research phase.

- State published price, discount, or rebate arrangements for recommended equipment and/or provide link to access company’s published prices for equipment and services.

- If the recommended equipment and related services are available for purchase on any existing Federal Supply Schedule contract(s) or other contracts against which NIST may be able to place orders, identify the contract number(s) and other relevant information.

- Identify any customers in the public or private sectors to which you provided the recommended or similar equipment.  Include customer(s) information: company name, phone number, point of contact, email address.

- Provide any other information that you believe would be valuable for NIST to know as part of its market research for this requirement.

- State if you require NIST to provide additional information to improve your understanding of the government’s requirement and/or would like to meet with NIST representatives to discuss the requirement and the capabilities of the identified equipment.

QUESTIONS REGARDING THIS NOTICE

Questions regarding this notice may be submitted via email to the Primary Point of Contact and the Secondary Point of Contact listed in this notice. Questions should be submitted so that they are received 5 days prior to the response date. If the Contracting Officer determines that providing a written amendment to this notice to document question(s) received would benefit other potential respondents, the questions would be anonymized, and a written response to such question(s) would be provided via an amendment to this notice.

IMPORTANT NOTES

This notice is for market research purposes and should not be construed as a commitment by NIST to issue a solicitation or ultimately award a contract.  There is no solicitation available at this time.

This notice does not obligate the Government to award a contract or otherwise pay for the information provided in response.

NIST reserves the right to use information provided by respondents for any purpose deemed necessary and legally appropriate.

Any organization responding to this notice should ensure that its response is complete and sufficiently detailed to allow the Government to determine the organization’s capability.

Respondents are advised that the Government is under no obligation to acknowledge receipt of the information received or provide feedback to respondents with respect to any information submitted.

After a review of the responses received, a synopsis and solicitation may be published on GSA’s eBuy or SAM.gov. However, responses to this notice will not be considered an adequate response to any such solicitation(s).

The responses shall not exceed 20 pages including all attachments, charts, etc.

Thank you for taking the time to submit a response to this request.